Beam Instrumentation – Part II
Ralph J. Steinhagen, CERN
Acknowledgements: R. Jones, P. Forck & U. Raich

---
Overview
- Beam instrumentation are the 'eyes and ears' of accelerators.
- Part I: electro-magnetic pick-ups (yesterday)
  - Beam intensities: Faraday-Cup, Fast-BCT, DC-BCT, WCM
  - Beam position: position and long. profile: Button-, Strip-line-, and Cavities
- Part II: beam loss and transverse beam profile
  - SEMs, wire scanner, OTR screens, luminescence, synch-light
  - Ionisation chamber, diodes, diamonds, scintillators, (cherenkov)
  - Special ultra-fast devices: electro-optical sampling & streak-camera
An accelerator can never be better than the instruments measuring its performance!	→ Important skill to assess whether beam observations are 'new/known physics', 'instrumental', or to guide whether/how performance can be improved.

---
Figure of Merit for most Accelerators
- For a collider – luminosity Lpeak:
- For a synchrotron-light source – Brilliance B:
{font fixed {L}}_{peak} ~sim~ {N_1 cdot N_2}over{A} ~sim~ {{N_1 cdot N_2}over{%pi %sigma_x %sigma_y}}
B ~sim {N_{photons}}over{A cdot d%phi d%theta} ~sim {{N_1~rightarrow~N_1^2}over{%pi %sigma_x %sigma"'"_x %sigma_y %sigma"'"_y}}
Increase this → intensity instrumentation
Decrease this → profile instrumentation
(todays focus)

---
The Typical Instruments
The Typical Instruments
- Beam Intensity
  - Faraday cups, beam current transformers, wall-current monitors
- Beam Position
  - electrostatic or electromagnetic pick-ups and related electronics
- Beam Profile
  - secondary emission grids and screens
  - wire scanners
  - synchrotron light monitors
  - ionisation and luminescence monitors
  - Femto-second diagnostics for ultra short bunches
- Beam Loss
  - Cherenkov optical fibre, ionisation chambers or pin diodes
- Machine Tune, Chromaticity and Luminosity → diagnostics tutorial

---
Secondary Emission (SEM) Grids
Secondary Emission (SEM) Grids
- When the beam passes through secondary electrons are ejected from the wires
- The liberated electrons are removed using a polarisation voltage
- The current flowing back onto the wires is measured
- One amplifier/ADC chain is used for each wire

---
Profiles from SEM grids
Profiles from SEM grids
- Charge density measured from each wire gives a projection of the beam profile in either horizontal or vertical plane
- Resolution is given by distance between wires
- Used only in low energy linacs and transfer lines as heating is too great for circulating beams

---
Wire Scanners
- A thin wire is moved across the beam
  - has to move fast to avoid excessive heating of the wire and/or beam loss
- Detection
  - Secondary particle shower detected outside the vacuum chamber using a scintillator/photo-multiplier assembly
- Secondary emission current detected as for SEM grids
  - Correlating wire position with detected signal gives the beam profile
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+

---
Beam Profile Monitoring using Screens
Beam Profile Monitoring using Screens
- Optical Transition Radiation
  - Radiation emitted when a charged particle beam goes through the interface of 2 media with different dielectric constants
  - surface phenomenon allows the use of very thin screens (~10mm)
OTR Screen
First full
LHC turn
10/9/2008
Exit window
Intensifier -
CCD
Un-Captured beam sweeps through he dump line
Mirror
Lens
Beam

---
Profile Monitoring using Screens
- Screen Types
  - Luminescence Screens
    - destructive (thick) but work during setting-up with low intensities
- Optical Transition Radiation (OTR) screens
  - much less destructive (thin) but require higher intensity

---
Light output from various Scintillating Screens
Example: Color CCD camera
  - very different light yield i.e. photons per ion‘s energy loss
  - different wavelength of emitted light
Al2O3
CsI:Tl
Al2O3:Cr
P43
YAG:Ce
Herasil
Quartz:Ce
ZrO2:Mg

---
Profile Monitoring using Screens
- Usual configuration
  - Combine several screens in one	housing e.g.
    - Al2O3 luminescent screen	for setting-up with low intensity
    - Thin (~10um) Ti OTR screen	for high intensity measurements
    - Carbon OTR screen	for very high intensity operation
- Advantages compared to SEM grids
  - allows analogue camera or CCD acquisition
  - gives two dimensional information
  - high resolution: ~ 400 x 300 = 120’000 pixels for a standard CCD
  - more economical
    - Simpler mechanics & readout electronics
  - time resolution depends on choice of image capture device
    - From CCD in video mode at 50Hz to Streak camera in the GHz range

---
Luminescence Profile Monitor
Luminescence or Rest-Gas-Ionisation
Profile Monitor
N2 excited state
Photon
emitted
e-
N2 ground state
Beam
Ion beam
Blackened walls
Vacuum gauge
Valve
Viewport
150mm flange
Lens, Image-Intensifier
and CCD Camera
N2-fluorescent gas
equally distributed

---
Luminescence or Rest-Gas-Ionisation
Profile Monitor
Beam Size
CERN-SPS Measurements
- Profile Collected every 20ms
- Local Pressure at ~510-7 Torr
2D
Side view
3D
Image
Time
Beam Size
Time
Injection
Beam size shrinks as
beam is accelerated
Fast extraction
Slow extraction

---
The Synchrotron Light Monitor
Beam
Synchrotron Light from
Bending Magnet
or Undulator

---
The Synchrotron Light Monitor

---
Measuring Short Bunch-Lengths
- Hadron accelerators:
  - σt ≈ 0.1 … 250 ns ↔ bunch lengths of few cm to tens of m (!!)	→ Electro-magnetic pick-ups, e.g. wall-current- or strip-line monitors
- Next Generation FELs & Linear Colliders
  - Use ultra short bunches to increase brightness or improve luminosity
- Q: How does one measure	such short bunches?
p+ @ PSB	~250 ns
p+ @ LHC	0.1-1 ns
H- @ SNS	100ps
e- @ ILC	500fs
e- @ CLIC	130fs
e- @ XFEL	80fs
e- @ LCLS	<75fs

---
Streak-Camera Principle
- Idea: a mini accelerator or inverse cathode ray tube (old TV)
  - typical resolution: σt~1 ps (Hamamatsu FESCA-200 σt=0.2 ps)

---
Streak Camera Realisation
acceleration
focusing
deflection
30 cm
Input	optics
Streak tube
CCD
PC
Fast Timing Generator
PC
Klystron

---
Streak Camera Results
APS kicker-induced synchro-betatron motion
courtesy Bingxin Yang, ANL
Slow Scan
Fast Scan
bunch length
SOLEIL: σt=35 ps

---
Electro-Optic Sampling – Non Destructive
Spectral Decoding
Limited to >250fs by laser bandwidth
Temporal decoding
Limited to >30fs by sampling laser pulse

---
Transverse Deflecting Cavity
- Working principle very similar to streak camera
  - ie. without photo luminescent screen causing a slow photo-response
  - destructive measurement → typ. placed in dedicated diagnostics sections
- State-of-the-art bunch length resolutions1: σt ≈ 1.3 fs
1Patrick Krejcik,”Commissioning the New LCLS X-band Transverse Deflecting Cavity with Femtosecond Resolution”, IBIC'13, http://www.ibic2013.org/prepress/talks/tual2_talk.pdf

---
The Typical Instruments
- Beam Intensity
  - Faraday cups, beam current transformers, wall-current monitors
- Beam Position
  - electrostatic or electromagnetic pick-ups and related electronics
- Beam Profile
  - secondary emission grids and screens
  - wire scanners
  - synchrotron light monitors
  - ionisation and luminescence monitors
  - Femto-second diagnostics for ultra short bunches
- Beam Loss
  - Cherenkov optical fibre, ionisation chambers or pin diodes
- Machine Tune, Chromaticity and Luminosity → diagnostics tutorial

---
Beam Loss Monitoring – Requirements I/II
Life isn't perfect so aren't accelerators
Machine Protection Protection:
  - interlock signal for fast (& safe) beam extraction to protect sensitive devices	(e.g. quenches in super-conducting magnets, de-magnetisation of undulators)
Beam diagnostics:
  - optimize transmission to the target.
  - Beam alignment to prevent unnecessary activation and complicated material handling → ALARA principle (As Low As Reasonably Achievable accelerator operation philosophy)
Accelerator physics:
  - using these sensitive particle detectors.
secondary beam
(accelerator, beam transfer line)
primary beam
(accelerator)
γ
e±,μ±
synchrotron
radiation
?
target
γ
e±,μ±
γ
e±,μ±
γ
e±,μ±
γ
e±,μ±
γ
e±,μ±
γ
e±,μ±

---
Energy Loss of Particles in Matter
Courtesy Lars Fröhlich

---
Beam Loss Monitoring – Requirements II/II
- Beam Loss Types:
  - Irregular or fast losses by malfunction of devices (magnets, cavities etc.)	→ BLM as online control of the accelerator functionality and interlock generation.
  - Intermediate
  - Slow regular losses
    - e.g. by lifetime limits or due to collimator → BLM used for alignment.
    - ALARA: As Low As Reasonably Achievable → minimise unnecessary activation of the machine complicating material handling
- Requirements for BLM:
  - High sensitivity to detect behaviour of beam halo e.g. at collimator
  - Large dynamic range:
    - low signal during normal operation, but very large signal in case of malfunction
    - detectable without changing the full-scale-range, e.g. scintillators from 102...107/s in counting mode.

---
The Large Hadron Collider LHC
The Large Hadron Collider LHC
Installed in the LEP tunnel, 27 km, Depth of 70-140 m
the Alps
Lake of Geneva
Sector 34
Jura mountains
LHC ring

---
27 km Circumference – 1232 LHC dipole magnet
B field 8.3 T (11.8 kA) @ 1.9 K (super-fluid Helium)
two-in-one magnet design:	two beam tubes with an opening of 56 mm (210 mm separation)
Operating challenges:
Very low quench levels (~ mJ/cm3) in an environment that stores MJ → GJ
Control of particle beam stability and losses is paramount!

---
Stored Beam Energies
Stored Beam Energies
x 200
XFEL (1s)
x 10 000
Quench Levels	Units	Tevatron	RHIC	HERA	LHC
Instant loss (0.01 - 10 ms)	[mJ/cm3]	4.5	18	2.1 - 6.6	87
Steady loss (> 100 s)	[mW/cm3]	75	75		5.3
[Speaker Notes: LHC extrapolation Energy in beam 7 times Tevatron Intensity of beam 30 times Tevatron Number of LHC of moving elements in the LHC about 200 Energy stored in both beams:	0.7 GJoule heat up from cryogenic temperature and melt 0.1 m3 of Copper]

---
Maximum LHC Energy of 7 TeV
LHC superconducting dipoles may loose superconducting state (“quench”)
minimum quench energy EMQE @7 TeV for t~10 – 20 ms
EMQE < 30 mJ/cm-3 vs. Estored = 350 MJ/beam
→ sufficient to quench all magnets and/or may cause serious damage
requires excellent control of particle losses
Example: un-controlled vs. controlled energy release
C = 5.4 1012 protons @ 450 GeV
D = 7.9 1012 protons @ 450 GeV
Vacuum pipe of QTRF in TT40
3.4 1013 protons @450 GeV
holes
for details see: Chamonix XIV:
 “Damage levels - Comparison of Experiment and simulation” and PAC'05
courtesy V. Kain

---
Beam Loss Detectors
Beam Loss Detectors
- Role of a BLM system:
  - Protect the machine from damage
  - Dump the beam to avoid magnet quenches (for SC magnets)
  - Diagnostic tool to improve the performance of the accelerator
- Common types of monitor
  - Long ionisation chamber (charge detection)
    - Up to several km of gas filled hollow coaxial cables
    - Position sensitivity achieved by comparing direct & reflected pulse
      - e.g. SLAC – 8m position resolution (30ns) over 3.5km cable length
    - Dynamic range of up to 104
  - Fibre optic monitors
    - Similar layout with electrical signals replaced by Cerenkov light

---
Beam Loss Detectors
- Common types of monitor (cont)
  - Short ionisation chamber (charge detection)
  - Typically gas filled with many metallic electrodes and kV bias
  - Speed limited by ion collection time - tens of microseconds
  - Dynamic range of up to 108
iin(t) + Iref
iin(t)
LHC

---
Beam Loss Detectors
- Common types of monitor (cont)
  - PIN photodiode (count detection)
    - Detect MIP crossing photodiodes
    - Count rate proportional to beam loss
    - Speed limited by integration time
    - Dynamic range of up to 109

---
Beam Loss Detectors – New Materials
Beam Loss Detectors – New Materials
- Diamond Detectors (CVD)
  - Fast & sensitive
  - Used in LHC to distinguish bunch by bunch losses
  - Investigations now ongoing to see if they can work in cryogenic conditions
Courtesy of E. Griesmayer

---
BLM Threshold Level Estimation
BLM Threshold Level Estimation

---
Summary
- Beam instrumentation are the 'eyes and ears' of accelerators.
- Part I: electro-magnetic pick-ups
  - Beam intensities: Faraday-Cup, Fast-BCT, DC-BCT, WCM
  - Beam position: position and long. profile: Button-, Strip-line-, and Cavities
- Part II: beam loss and transverse beam profile
  - SEMs, wire scanner, OTR screens, luminescence, synch-light
  - Special ultra-fast devices: electro-optical sampling & streak-camera
  - Cherenkov fibre, Ionisation chamber, diodes, diamonds, scintillators
An accelerator can never be better than the instruments measuring its performance!	→ Important skill to assess whether beam observations are 'new/known physics', 'instrumental', or to guide whether/how performance can be improved.

---
Quench levels
Quench levels