Feasibility of (semi-) automated Tune Control in the PS
BI-OP Meeting on PSB/CPS Tune Measurement, 2007-09-25
Top-Level Base-Band-Tune (BBQ) Measurement Functionalities
as seen from the OP GUI
Stephane Bart Petersen, Andrea Boccardi, Marek Gasior, Lars Jensen, Rhodri Jones, Peter Karlsson, Ralph Steinhagen

---
Outline
FESA interface properties:
  - 'Settings'
  - 'Acquisition' ('Measurements')
Expected performance:
  - analogue BBQ front-end
  - digital acquisition board (DAB) and FESA server
Why “operational knowledge” is required for a robust tune reconstruction
  - other-than-tune peaks, betatron coupling, chromaticity,	fast longitudinal bunch operations/length variations, ...

---
FESA Interface for the BBQ DAB
As seen from operation:
  - FESA server names: PSB: 'BR.BQSB', PS: 'PR.BQSB', SPS: 'SPSBQSB'
  - Main properties:
    - 'Settings': basic acquisition settings
    - 'Acquisition': the results
    - 'Measurement': as 'Acquisition' but on	a sample-by-sample basis	(used in case of coasting beam)
Naming: 'one measurement' = 'one FFT spectra based on <nbofAcq> turns'
  - after post-processing: → one Q1, Q2, Qx, Qy, |C-| measurement
Further details accessible through the FESA navigator:
  - http://wwwpsco.cern.ch/private/java/fesa/CURRENT_RELEASE/FNT.jnlp
  - or 'Tune Viewer' example application (all CERN accelerators):http://slwww/~pcrops/releaseinfo/pcropsdist/sps/sps-tuneviewer/PRO/

---
'Settings' Property I/VI
Acquisition trigger settings:
acqState: <on,off>:
  - enables or disables the acquisition
  - default: 'off', also switches 'off' once 'single' acquisition is finished
acqMode: <SINGLE, CONTINUOUS>
  - 'SINGLE': perform 'nbOfAcq' measurements and then switch to state 'acqState=off' (fail-safe mechanism)
  - 'CONTINUOUS': 'coast mode' (ignores cycle alignment)
externalStartTrigger: <true, false>
  - switches 'start first acquisition' between external HW based (BST-Master) trigger or CTRP timing card (e.g. 'cycle start') trigger
  - required to synchronise with external events such as collimator movements or RF that are not aligned to CERN's UTC 'ms' clock.
  - default: 'false'

---
'Settings' Property II/VI
Acquisition trigger settings:
nbOfTurns: <T256, T512, T1024, ..., T262144>
  - length of acquired data window(s) in turns.
  - default: CPS/PSB/SPS: 'T1024'; LHC: 'T8192'
nbOfMeas: <int>
  - number of requested measurements (N.B. '0' → '1')
  - default: '1' (maximum limited by available memory)
acqOffset: <int>
  - offset of the first measurements in 'ms'
  - default: CPS/PSB: '10', SPS: '25', LHC: '1000'
acqPeriod: <int>
  - distance in between measurements in 'ms'
  - special mode 'acqPeriod == 0': 'back-to-back' acquisition (BI experts/MD type acquisition)
  - default: CPS/PSB: '10', SPS: '25', LHC: '1000'

---
'Settings' Property III/VI
Excitation trigger settings:
exDelay: <int>
  - number of measurements to perform before first excitation [acqPeriod]
  - default: '0'
exOffset: <int>
  - turn delay of excitation (kick) w.r.t. to individual acquisition start
  - default: '0'
exPeriod: <int>
  - number of non-excited measurements (gap) in between measurements with excitations. e.g. excitationGap = 3 → every fourth measurement contains excitations
  - unit: [acqFrequency]
  - default: '0'

---
'Settings' Property IV/VI
Data pre- and post-processing:
dataMode: <RAW, FFT, MAG, RAW_AND_FFT>
  - RAW: raw oscillation data (N.B. the BBQ ≠ a BPM: oscillation frequencies below 0.1frev are attenuated by a significant amount)
  - FFT: Fast-Fourier-Transform spectrum (contains real- and imaginary part)
  - MAG: magnitude spectrum (no phase information)
  - default: PSB/CPS: 'MAG', SPS/LHC: 'RAW_AND_FFT'
windowFunction: <RECTANGULAR, HAMMING, HANN, BLACKMAN, NUTTALL, BLACKMAN-HARRIS, BLACKMAN-NUTALL, FLAT_TOP>
  - apodisation function of the FFT routine (N.B. the 'RAW' data is unaffected by this)
  - default: 'HANN'

---
'Settings' Property V/VI
General excitation settings:
exMode: <NOEXCITATION, CHIRP, KICK>
  - chooses the excitation source
  - default: 'NOEXCITATION'
    - Lots of diagnostic possibilities (tune, chromaticitiy, RF problems, instabilities ....). However: no guarantee that the tune is always visible without excitation
enableExcitationH, enableExcitationV: <true/false>
  - default: 'false' (fail-safe)
  - N.B. <true, true> → forces alternating of chirp in H/V plane
Chirp related settings:
exAmplitudeH, exAmplitudeV: <float> [defined by chirp/kick exciter]
  - N.B. hardware (kicker) to be modified for 2008 startup
chirpStartFreqH, chirpStopFreqH, chirpStartFreqV, chirpStopFreqV: <float>
  - start/stop frequency ([frev]) of chirp excitation (usually: ]0.0, 0.5])

---
'Settings' Property VI/VI - Summary
start of cycle
acqOffset [ms]
nbOfMeas []
...
acqPeriod [ms]
nbOfTurns []
exPeriod [acqPeriod]
exDelay [acqPeriod]
enableExcitationH/V
exAmplitudeH/V
...
exOffset [turns]
Common measurement scenario:
  - measure (acqState=on) from injection (acqOffset=0) every 10 ms (acqPeriod=10) till 900 ms (nbOfMeas=floor(900/acqPeriod)) while kicking the beam every measurement (exDelay=0) in the vertical plane (exMode=KICK, enableExcitationH = true, enableExcitationV = false)
Example application: 'Tune Viewer' :http://slwww/~pcrops/releaseinfo/pcropsdist/sps/sps-tuneviewer/PRO/
measurement
excitation

---
DAB based BBQ acquisition: PS examples
PS-MD2, H/V kicks

---
DAB based BBQ acquisition: PS examples
SFTPRO, H kicks only,
Simple peak detection: “highest peak” or “highest S/N ratio”

---
Guaranteed Performance and Robust Tune Detection
or: ... “not all that glitters is gold”
DAB/FESA provides for the PS/PSB:
  - transverse FFT spectra based on 1024 turns every 5 ms for one nominal CPS cycle (1.2 seconds)
    - both planes! (old system: either H/V@5ms or H&V@10ms)
    - memory bandwidth limit: 'acqPeriod'·'nbOfTurs' = const.)
  - some cycle do not require excitation. However, there is no guarantee for that (monitor only)
“Highest peak” or “highest S/N ratio” does not guarantee that only the tunes are detected, robust detection depends on:
  - operational tune range and working point (Q < 0.5 v Q > 0.5)
  - expected tune drifts
  - whether cycle requires full or no betatron-coupling
  - knowledge of RF operation
  - a rough estimate of the machine chromaticity (→ to steer excitation strength)

---
PS-SFTPRO cycle, back-to-back acquisition, H/V kicks every 5 ms, horizontal plane, “back-to-back”
time in cycle
tune

---
PS-SFTPRO cycle, back-to-back acquisition, H/V kicks every 5 ms, vertical plane
time in cycle
tune

---
Betatron-Coupling: SFTPRO1 I/II

---
Betatron-Coupling: SFTPRO1 II/II
after reconstruction
Tracking of eigen-modes:

---
PS-SFTPRO cycle
Betatron-Coupling after Injection

---
Preliminary Conclusions
Our target:
  - same (/similar) low-level front-end electronic, the same digital acquisition electronic for all CERN accelerator
  - > 80% of common DAB/FESA functionality
    - the 'missing 20%' likely used during MDs (all CERN) and LHC
The robust tune reconstruction for all possible operational machine and beam condition requires estimates and previous knowledge on
  - operational tune range and working point (Q < 0.5 v Q > 0.5)
  - expected tune drifts
  - whether cycle requires full or no betatron-coupling
  - a rough estimate of the machine chromaticity (→ to steer excitation strength)

---
additional supporting slides

---
PS-SFTPRO(1) cycle

---
PSB - SFTPRO
four-times oversamples ('tune =1.0' = 4 frev)

---
Tune Measurement with slowly-extracted un-bunched Beam in the SPS
chirp off
RF + damper
noise
chirp on
RF off
vertical tune
horizontal tune
second injection
damper noise
first injection

---
Q/ξ Beam Response Matrix Uncertainties
Uncertainties in the beam response matrix reduced the effective control/feedback bandwidth but does not affect the steady-state precision
E.g. LHC orbit feedback:
perfect optic	→ 1 iteration
20% beta-beat	→ ~2 iterations
20% calibration error	→ ~7 iterations

---
Non-linear Slew-rate Limited Exciter Response
LHC orbit dipole corrector: ΔI=0.01 ↔ Δx ≈ 15 μm @7TeV

---
Cross-Dependability and Constrains of FB Loops II/III - Coupling I/II
Strictly speaking: PLL measures eigenmodes (Q1, Q2) which in the presence of coupling may be rotated w.r.t. unperturbed tunes (qx, qy, Δ = |qy – qy|):
Possible improvement:
  - optimise tune working point (larger tune-split),
  - vertical orbit stabilisation in lattice sextupoles,
  - active compensation and correction of coupling
Q_{1,2} = {1 over 2} left( q_x + q_y +- sqrt{%DELTA^2 + {abs{C^{-%}}}^2} right)
Δ
|C-|
Tune control on Q1,Q2 only
would break here
qy
qx
Q2
Q1
courtesy P. Cameron, BNL
RHIC, 2005

---
Cross-Dependability and Constrains of FB Loops II/III - Coupling II/II
Measure ratio between regular and cross-term:
  - A1,x: “horizontal” eigenmode in vertical plane
  - A1,y: “horizontal” eigenmode in horizontal plane
Decoupled feedback control
  - qx, qy	→	quadrupole circuits strength
  - |C-|, χ	→	skew-quadrupole circuits strength
color red {r_1 = A_{1,y} over A_{1,x}} ~~and~~ color red {r_2 = A_{2,x} over A_{2,y}} newline newline alignl % drarrow ~~~ abs{C^{-%}} = abs{Q_1 - Q_2} cdot {{2 sqrt{r_1 r_2}} over {(1 + r_1 r_2)}} ~~and~~ %DELTA = abs{Q_1 - Q_2} cdot {{(1 - r_1 r_2)} over {(1 + r_1 r_2)}}
χ
implemented and tested at RHIC
R. Jones e.al., “Towards a Robust Phase Locked Loop Tune Feedback System”, DIPAC'05, Lyon, France, 2005
P. Cameron, this workshop's Poster Session

---
PS magnet - old layout
Controllability -
PS' Combined Function Magnets
Main Dipole's Pole-Face-Winding (PFW) Schematic (before 1978)
1x Focus. (FW, FN), 1x Defocus. (DW, DN), 8-loop (~octupole) → 3 circuits
narrow
wide
2x Focusing (
2x Defocusing

---
HOT: SPS Q' tracking study
Radial RF modulation dp/p=1.6∙10-5, set ξ := 0.05
ΔQmod≈ 10-5