Title
Some Notes on:
Continuous LHC Beta-Beat Measurements
A. Boccardi, E. Calvo Giraldo, M. Gasior,
J.L. Gonzales, R.Jones, R.J. Steinhagen
Accelerator and Beams Department
Beam Instrumentation Group

---
Motivation
- Motivation: System dependence on known/constant beta-function:
  - Machine Protection and Collimation, Physics, Squeeze Diagnostics
  - Classic methods: 'kick'-type excitation & BPMs, K-modulation & Q-PLL, Closed-orbit-response (LOCO)	→ cannot achieve the required precision/time-scales under nominal conditions!
- β-Phase Advance Method - Turn-by-Turn
  - BBQ based Test-Setup in SPS LSS5
  - Systematic and statistical noise contribution
  - Exploitation Examples: SPS lattice drifts & off-momentum beta-beat
- β-Phase Advance Method – Orbit → Report in Preparation
- Next Steps & Control of Betatron-Function:

---
Compromised Machine Protection via Orbit Bumps
Combined failure: Local orbit bump and collimation efficiency (/kicker failure):
- Primary collimator (TCP) limits |xb(s)|max locally to <5.7s, secondary collimator (TCS) at~ 6.7s
- To guarantee two stage cleaning efficiency/machine protection:
  - Local: TCP must be >0.7s closer than TCS w.r.t. the beam → Orbit FB
  - Global: no other object (except TCP) closer to beam than TCS
→ Orbit bumps may compromise function of machine protection/collimation
→ tackled by LHC Orbit Feedback
IR2
IR7
e.g 'bump in arc'
Potentially:
< 6.7s
Na [σ]
~7.5σ
primary halo
6.7s
5.7s
closed orbit
TCP & TCS
MKI
TDI

---
Compromised Machine Protection through Beta-Beat
Combined failure: beta-beat and collimation efficiency
“Collimator gap must be 10 times smaller than available triplet aperture!”1
  1. β-Beat reduces required protection: Δβ/β ≈ 20 % → 20% tighter collimator settings
  2. β-Beat reduces cleaning performance
IR2
IR7
triplets
A) Potentially
< 6.7s
B) “Secondary” collimator
becomes primary (here: Δβ/β = 1)
Na [σ]
~7.5σ
6.7s
5.7s
closed orbit
TCP & TCS
MKI
TDI
~ 0.15
~ 0.6
a_{coll}~leslant~ a_{triplet} `cdot` sqrt{%beta_{coll} over %beta_{triplet}} `cdot` left({A^{max}_{primary}} over {A^{max}_{secondary}}right)
1 R. Assmann, “Collimation and Cleaning: Could this limit the LHC Performance?”, Chamonix XII, 2003

---
Performance Limitations & Constraints on β*
N.B. C = βtrip· β*
β-beat meas. +
RT-Feedback?
IR Upgrade
LHC Orbit Feedback

---
Collimation Performance Limitations & Constraints on β*
Maximum allowed safe beam intensity1,2:
  - Min. accept. lifetime:	Τmin ≈ 10 min.
  - Dilution length:	Ldil ≈ 50 m
  - Quench level (@7 TeV) Rq:	Rq ≈ 7.6∙106 prot./m/s
  - Collimation inefficiency:	η
Peak-Luminosity:
color red {N_{max}} leslant {%tau_{min} cdot R_q cdot L_{dil.}} over %eta
L_{max} approx {{1} over {4 %pi}} cdot {{color red {N_{max}} cdot n_b f_{rev}} over {%beta^{% * % } %epsilon}}
Collimation inefficiency vs. β-beat (β*=0.55 m)
MAC Dec 2004
nominal:
Δx≈35 μm@coll
need to operate here!
courtesy R. Assmann
1 R. Assmann, “Collimation and Cleaning: Could this limit the LHC Performance?”, Chamonix XII, 2003
2 S. Redaelli, “LHC aperture and commissioning of the Collimation System”, Chamonix XIV, 2005
3 R. Steinhagen, “Closed Orbit and Protection”, MPWG #53, 2005-12-16
Coll. system
version ~ 2002
Collimation inefficiency vs. orbit error
courtesy R. Assmann
peak-to-peak orbit error [σ]

---
Constraint II: Diffractive Physics
p-p Luminosity using Optical Theorem
Special parallel to point focusing machine optic (b0 ≈ 1600 m)
  - Roman Pots move close to the beam halo, measure dN/dt down to:
  - Observables: abs. Luminosity, total p-p cross-section, diffractive physics
    - Requires good knowledge on
      - Beta-functions b0 at IP and bd at detector
      - Beam momentum p
      - minimum distance of roman pot xmin w.r.t. beam centre
  - Desired: DL/L ≈1% → Dt/t ≈1% → 0.5∙Dq/q ≈ Dx/x ≈ 5∙10-3
    → absolute beam position stability at roman pot (xmin~ 1mm) < 5 mm!!
    → value of betatron function at IP and RP: → Dβ/β ≈1% !!
Roman pot with silicon tracker measures x:
quadrupole
scattered p
x~q
q
p
p
t_{min} = (p %theta_{min})^2 sim {p^2} over {%beta_0 %beta_d} cdot x_{min}^2

---
Limitations on Squeeze Diagnostic
Squeeze involves > 45 individual magnetic strength settings (Optics), so far: no continuous check on effective optics during/at the end of individual steps
- “Classic” methods may not reach/be compatible with nominal requirements
  - K-modulation induced Q-Changes:
    - Limit: knowledge on quadrupole transfer function	(hysteresis, D&S, β|max ≈ 4.2km & ΔQmax<10-3 → Δk/knom<5·10-5)
  - Kick + turn-by-turn analysis of BPM (phase and/or amplitude), limits:
    - Potential particle loss (beta-functions at triplet) & emittance blow-up
    - Systematic phase errors, amplitude detuning/Landau damping
      - large kicks may probe phase advances (dynamic aperture) which may not be representative for nominal beam operation
      - beam will be collimated at 6 sigma (kick amplitudes < 1.2 mm @7TeV)!
    - … not ideal for continuous monitoring/regular operation (ε blow-up).
  - Closed orbit response analysis (LOCO):
    - resolution/performance compatible with nominal operation
    - Limit: scan requires several minutes per IP (full scan: ~2 OP-shifts)
%DELTA Q ~approx~ {1 over {4%pi}} cdot %beta(s) cdot %DELTA k(s)

---
Betatron Function via Phase-Advance Measurement
- Long history at CERN. Original idea dates back to AB-BI report (doctoral thesis)	P.Castro, Luminosity and Betatron Function Measurement at [..] LEP, CERN SL/96-70 (BI)
- … beating in amplitude related to beating in phase:
Phase sampling:
Beta-Beat reconstruction (FB/Control would work with phases):
{{%DELTA%beta} over %beta}(s) ~=~ {{1} over {2`sin(2%pi Q)}} lint { %beta_k } cos(2 cdot abs {%mu(s)- %mu(a)} - 2 %pi Q) %DELTA k(a)`da
%mu(s) := int csub {0} csup{s} {1 over {%beta(a)}} da
{{%DELTA %mu} over {%mu}} (s) ~ sim ~ {{%DELTA %beta} over {%beta}} (s)
Δφ
-Δφ
Case I:
Δμ12 =45° - Δφ, Δμ13 =90° - 2 Δφ
0
Δφ
Case II:
0
0
Δμ12 =45° + Δφ, Δμ13 =90°
{%DELTA %beta_1} over {%beta_1} ~=~ {cot (%DELTA%mu_{12}^{meas.}) `-` cot (%DELTA%mu_{13}^{meas.})} over {cot (%DELTA%mu_{12}^{theo.}) `-` cot (%DELTA%mu_{13}^{theo.})}
{%DELTA %beta_2} over {%beta_2} ~=~ {cot (%DELTA%mu_{12}^{meas.}) `-` cot (%DELTA%mu_{23}^{meas.})} over {cot (%DELTA%mu_{12}^{theo.}) `-` cot (%DELTA%mu_{23}^{theo.})}
{%DELTA %beta_3} over {%beta_3} ~=~ {cot (%DELTA%mu_{23}^{meas.}) `-` cot (%DELTA%mu_{13}^{meas.})} over {cot (%DELTA%mu_{23}^{theo.}) `-` cot (%DELTA%mu_{13}^{theo.})}
N.B. Phase-Beating usually used for correction!
45°
45°

---
Beta-Beat/LHC BPM Prototype System in the SPS-LSS5
βv≈100 m
βv≈100 m
βv≈100 m
βv≈20 m
βv≈100 m
βv≈100 m
Measurement (markers), sinusoidal fit (solid line):
BPMB.513
BPMB.515
BPMB.517
BPMB.518
BPMB.519
BPMB.520
{%DELTA %beta_1} over {%beta_1} ~=~ {cot (%DELTA%mu_{12}^{meas.}) `-` cot (%DELTA%mu_{13}^{meas.})} over {cot (%DELTA%mu_{12}^{theo.}) `-` cot (%DELTA%mu_{13}^{theo.})}
{%DELTA %beta_2} over {%beta_2} ~=~ {cot (%DELTA%mu_{12}^{meas.}) `-` cot (%DELTA%mu_{23}^{meas.})} over {cot (%DELTA%mu_{12}^{theo.}) `-` cot (%DELTA%mu_{23}^{theo.})}
{%DELTA %beta_3} over {%beta_3} ~=~ {cot (%DELTA%mu_{23}^{meas.}) `-` cot (%DELTA%mu_{13}^{meas.})} over {cot (%DELTA%mu_{23}^{theo.}) `-` cot (%DELTA%mu_{13}^{theo.})}
1P.Castro, Luminosity and Betatron Function Measurement at [..] LEP, CERN SL/96-70 (BI)
BPMB.517
BPMB.518
BPMB.519
Δμ12
Δμ13
beam
Δμ ≈ 90°
Δμ ≈ 90°
Δμ≈45°
Δμ≈45°
Δμ≈45°

---
Test Setup
- Yet another exploitation of the BBQ Principle:
  - AC-coupled peak detector:
    - no saturation, self-triggered, no gain changes between pilot and nominal
    - intrinsically down samples spectra: ... 6 GHz → 1kHz ... frev
    - Base-band operation: very high sensitivity/resolution ADC available
    - Measured resolution estimate: < 10 nm	→ ε blow-up is a non-issue
- Digital acquisition: HP Proliant 16”,	1U + M-AUDIO Delta 1010
  - 8 analogue inputs/outputs, 16”, 1U
  - frequency response: 20Hz-22kHz, +/-0.3dB
  - >100 dB dynamic range/S/N ratio
  - THD: 0.00072% (A/D), 0.00200% (D/A)
~1300CHF
630CHF
BBQ-FE
BPM
30++ dB isolation
in between ports
3dB/splitter
~65 CHF
DAB64x
DAB64x
DAB64x
WBTN-FE
DAB64x

---
Impact of Sharing the same Button Pick-up - Disclaimer
Tests with beam in the SPS indicate that there is no obvious cross-talk in between the regular LHC WBTN (orbit) and the tested diode-based acquisition electronic used for the continuous beta-beat measurement.
- However: sharing intrinsically halves signals seen by the acquisition chain
  - Reduced minimum intensity detectable by button-type BPM (2 → 4·109)
    - Only relevant for IR7, may be less of an issue:
      - redundancies in IR7: multiple BPMs per cell & collimator
      - only affected for below pilot intensities (< 4·109 p/bunch)
    - N.B. Not an issue for strip-line pickups and/or nominal beam:
      - signals are attenuated simplify intensity gain-switching
  - Ongoing tests in the lab...
If this proves to be really an issue, installation of additional pick-ups in the region of interest may be required.

---
Sample-by-Sample Phase Demodulation Scheme
Modified mixing scheme:
- Alternative to mixing method: IIR Hilbert transformer
  - trade-off: higher bandwidth ↔ lower phase precision
sin(2πfe+φw)
X
z1(t)
X
φw+phase
correction
R1(fe)∙sin(2πfe-φ1)
-
-
cos(2πfe+φw)
mod 2π
Σ
Σ
Δμ1x
sin(2πfe+φw)
+
rect2polar
Rect.
2
Pol.
X
R(ω)
%DELTA z_{1} {size 10 {(t)}} `approx` sqrt{%beta_k %beta_1} `cdot` sin(%omega_e t + %varphi_w + color red {%mu_1}) newline newline newline %DELTA z_{x} {size 10 {(t)}} `approx` sqrt{%beta_k %beta_x} `cdot` sin(%omega_e t + %varphi_w + color red {%mu_x}) newline
φx
zx(t)
I
φ
X
Rx(fe)∙sin(2πfe-φx)
Q
cos(2πfe+φw)
Rect.
2
Pol.
R(ω)
φ
φ1
I
Q
rect2polar
fLP
BP
BP

---
QE.603/QE.604 induced β-Phase-Advance Beating
QE.604 (β = 20m!!) – trims:
0A → +50 A → -50 A → 0 A
QE.603 (β = 103m) – trims:
0A → +50 A→ -50 A → 0 A

---
QE.603/QE.604 induced β-Beating
{%DELTA %beta_1} over {%beta_1} ~=~ {cot (%DELTA%mu_{12}^{meas.}) `-` cot (%DELTA%mu_{13}^{meas.})} over {cot (%DELTA%mu_{12}^{theo.}) `-` cot (%DELTA%mu_{13}^{theo.})}
Corresponding beta-beat:
- Measured beta-beat is compatible with magnet calibration curves.
- Peak-to-peak β-beat “noise”: ~ 0.5 %
  - unlikely due to diagnostic
  - seen already at LEP:	(though not time resolved)
 → real drift of the optics!
Δφtheo.= 0.08
vs. Δφmeas ≈ 0.2
P. Castro, CERN SL/96-70 (BI), p. 61

---
Beta-Beat Sensitivity and Error Estimates
IR optic
Residual resolution/systematic error
  - ARC optics: requires error below ~1°
  - IP optics: requires error below ~0.02°
N.B. Plots have logarithmic z-scale!
{%DELTA %beta_1} over {%beta_1} ~=~ {cot (%DELTA%mu_{12}^{meas.}) `-` cot (%DELTA%mu_{13}^{meas.})} over {cot (%DELTA%mu_{12}^{theo.}) `-` cot (%DELTA%mu_{13}^{theo.})}
arc optic
%DELTA%mu_{1i}^{meas.} := %DELTA%mu_{1i}^{theo.} + %DELTA%varphi_{1i}
45° optic
Δφ12≈ 176° (CMS/ATLAS)
operational range

---
Beta-Beat Measurement Error Sources I/II
Systematic Phase Errors
- Sources – usually depend on observation/excitation frequency
  - Systematic delays:
    - Pick-up to acquisition system cable length (e.g. 100 m@ QAC=0.25 frev)
      - SPS: Δφ ≈ 2° LHC Δφ ≈ 0.5°: Δβ/βsys.≈ 3-10% (45° lattice)
        → cable delay compensation mandatory for direct β*-Measurements
  - Low-frequency pre-processing and analogue front-end asymmetry (mostly filters, N.B. Current has been not optimised for those issues)
    - Delta 1010 – analogue pre-filter:	Δφ ≈ 7° (measured)
    - BBQ front-end: Δφ ≈ 10° (measured, here: only Chebychev stage shown)
    - Systematic drift: < 0.1° → will be further reduced
%DELTA%varphi [deg] ~=~ 360° cdot {{%DELTA%tau}{f}}
Qac
frev

---
Beta-Beat Measurement Error Sources II/II
Statistical Phase Noise
- Statistical noise adds vectorial to the carrier signal:
  - excitation amplitude (carrier signal):	A
  - noise in time (frequency) domain: σt (σf)
  - Equivalent number of turns:	N
matrix{ %sigma(%varphi) # %=% # alignl { arcsin left({%sigma_{f} over A}right)} # %=% # alignl {arcsin left(sqrt{2 over N}{%sigma_{t} over A}right)} ## % # % % approx % # alignl {sqrt{2 over N}{%sigma_{t} over A}} # % # % }
for small noise
to signal ratios
Delta 1010 intrinsic phase noise@1Hz
Best case BPM+kick:
S/N=1mm/0.2mm &
1024 turns
σ(φ)
A
σf
φ

---
Typical SPS Beam Spectrum
single bunch, ~ 7 ·1010 protons@270GeV (coasting)
ZOOM
- Based on 128k turns (~ 1.3 seconds)
  - noise floor (time domain): ~100 nm (BBQ1: ~5 nm)
  - driven 'AC-dipole' signal: ~20-30 um
  - LHC BPMs give ~30 dB less signal than BBQ1 installation (buttons vs. 30 cm strip-line)
  - Residual tune signals ~ 0.5/2 um (calibrated w.r.t. Signal seen on SPS BPMs)
  - off-resonance excitation → no emittance blow-up
driving frequency:
QAC ≈ 0.25
(RF damper exciter)
Qy ≈ 0.17
Qx ≈ 0.13

---
Residual Beta-Beat Drifts - Revisited
- Residual phase motion (blue: BPM1->2, red: 0.5*BPM1->3)
  - Acquisition/electronic induced	noise would be “equal”/randomly	distributed over all channels
  - GM induced sextupole shifts
    - Δx ≈ 100 um r.m.s.	→ Δβ/β ≈ 0.1-0.2%
    - bit too large to be the	only perturbation source...
Δφ
-Δφ
Case I:
Δμ12 =45° - Δφ, Δμ12 =90° - 2 Δφ
0
Δφ
Case II:
0
0
Δμ12 =45° + Δφ, Δμ12 =90°
45°
45°

---
Further Exploitation Possibilities
System can be further exploited for fast and transparent measurements of physics affecting Δβ/β that earlier required significant amount of beam time
- Example: vertical off-momentum β-Beat:
  - Continuous radial modulation: Δp/p ≈ 1·10-3 @ 1 Hz
  - One full measurement data set every second!
  - N.B. Step in phase → off-centre horizontal orbit in lattice sextupoles

---
Example: SPS LHC1 Cycle-to-Cycle Stability II/II
Phase-Beating
In between two coasts...
LHC1
injection
plateau
injection
plateau
LHC1

---
Example: SPS LHC1 Cycle-to-Cycle Stability II/II
- Beta-Beat
In between two coasts...
LHC1
injection
plateau
injection
plateau
LHC1

---
Next Steps I/II:
Improvement to LHC Test Installations:
- Beside the intrinsic loss of signal due to the 3dB-signal splitter, initial tests show that sharing and cross-talk effects in been the regular WBTN and Beta-Beat system appears to be minimal.
  - tested with 75 ns/50 ns/25 ns/single nominal & pilot bunch configurations
  - 3dB-Splitter are being further tested in the lab using a bunch simulator setup
  - Affects mainly performance with ultra-low intensity bunches (<2·109 p/bunch)
- Further Setup Improvement:
  - systematic drifts of analogue front-end stages 0.1°/0.5 hour (β* meas)
  - Scalability and possible system integration (in view of LHC application)
  - Install 3 (+2) β-beat acquisition chains, both planes, either B1 or B2, in parallel to the regular BPM system, e.g. in LHC beta-cleaning insertion:
    - the minimum installation without pulling additional cables
primary collimator

---
Next Steps II/II:
Beta-beat Control
The LHC Prototype system's usefulness is two-fold:
- Provided β does not change: study real beta-beat as a function of time and use measured values to possibly relax collimation requirements
- Provided β does change: same as above but – in addition –	use real beta-beat values as an input to a real-time feedback loop	(e.g. primary/secondary collimators, IPs)
  - Correction scheme similar to LHC Orbit FB system using the dispersion suppressor's and other individually powered quadrupoles	(N.B. we are not as “free” in correcting μ as for correcting the orbit)
  - Additional regions of interest: experimental insertions, Inj./Extr. IRs, ...
- The possible merit...
 → remove/reduce protection/cleaning limitations on β* & stored intensity
 → +20% effective Luminosity
        “Isn't this worth being further investigated?”
vec {{%DELTA %mu} over {%mu}} ~=~ {underline {R}}_{%mu} `cdot` vec{%delta}_{DS} ~~ rightarrow csup{SVD} ~~ vec{%delta}_{DS} ~=~ {widetilde {underline {R}}}_{%mu}^{-1} `cdot` vec {{%DELTA %mu} over {%mu}}

---
Conclusions
- A real-time β-beat measurement system has been successfully tested at the SPS based on the continuous measurement of the cell-to-cell phase advance.
  - Achieved resolution: Δβ/β < 1% @ 1 Hz measurement bandwidth
  - Required S/N ratio: ~ 20 μm/100 nm → ε blow-up is a non-issue
    - compatible with nominal LHC operation
  - Shared pickup scheme (to 1st order) compatible with regular WBTN oper.
    - price of sharing: loose ~ 3dB signal, OK for strip-line, buttons → may affect smallest detectable bunch intensity for some IR7 BPMs
Measured residual 1% drift of SPS lattice and off-momentum beta-beat
- Present limitations of the system:
  - For the time being: only local measurements
  - 45° optics (LHC arcs):	< 0.01° ↔ Δβ/β << 1% @ 1Hz bandwidth
    - residual beta-function stability and S/N ratio
  - Exp. Insertion optics:	~ 0.1° ↔ Δβ/β ≈ 30% (Δμ12 ≈ 178°)
    - systematic phase and drifts, can be improved, target: Δβ/βIP ≈ 3%
Provided LHCPerfC approval: We would like to install a similar setup in LHC-IR7 for 2009 to investigate residual beta-function drifts and possible exploitation of this measurement techniques within a real-time β-beat FB loop.

---
additional supporting slides

---
Example: SPS LHC1 cycle-to-cycle stability
necessary correction
Fourier Spectrum before and after band-pass filter (carrier at 10.8 kHz)
before band-pass filtering
after band-pass filtering
Qac
Qy
Qac
Qx

---
LHC Q Base-Line Q Instrumentation
Back-bone: Base-Band-Q Principle on a Slide
BBQ detector
BBQ front-end
- Basic principle: AC-coupled peak detector
  - no saturation, self-triggered, no gain changes between pilot and nominal
  - intrinsically down samples spectra: ... 6 GHz → 1kHz ... frev
    - Base-band operation: very high sensitivity/resolution ADC available
    - Measured resolution estimate: < 10 nm → ε blow-up is a non-issue
- Since 2007: added low-pass filter in front of diodes to suppress intrinsic longitudinal (carrier) signal propagation into transverse oscillations

---
Beta-Beat Sources
- Beta-Beat Sources
  - Quadrupole gradient and Momentum errors
  - Feed-down due to off-centre horizontal orbit in lattice sextupoles
Requirements: Brüning, Fartoukh, LHC Project Report 501

---
Space Domain: - No “black feedback magic”
Effects on orbit, Energy, Tune, Q' and C- can essentially cast into matrices:
  - LHC matrices' dimensions:
  - control consists essentially in inverting these matrices:
- Some potential complications:
  - Singularities = over/under-constraint matrices, noise, element failures, spurious BPM offsets, calibrations, ...
  - Time dependence of total control loop → “The world goes SVD....”
%DELTA vec{x}(t) = underline R cdot vec%delta (t) ~with~ R_{ij} = {alignc sqrt{%beta_i %beta_j}} over {2 sin(%pi Q)} cdot cos(%DELTA %mu_{ij}-%pi Q) + {D_i D_j} over {C(%alpha_c-1/%gamma^2)}
matrix multiplication
{underline R}_{orbit} in setR^{1070 times 530}
{underline R}_{Q} in setR^{2 times 16}
{{underline R}_{C^{-% }} in setR^{2 times 10/12}}
{underline R}_{Q'} in setR^{2 times 32}
color black { {left ldline {vec x}_{ref} - {vec x}_{actual} right rdline}_2 = {left ldline underline R cdot {vec %delta_{ss}} right rdline}_2 < %epsilon ~ rightarrow~ vec %delta_{ss} = {tilde R}^{-1} %DELTA vec{x} }

---
SVD on a slide
Space-Domain:
Singular Value Decomposition (SVD) on a slide
Linear algebra theorem*:
eigen-vector relation:
T
color black {alignl %lambda_i {vec u}_i = {underline R} cdot {vec v}_i} newline color black {alignl %lambda_i {vec v}_i = {underline R}^T cdot {vec u}_i}
color white { dlrarrow }
though decomposition is numerically more complex final correction is a	simple vector-matrix multiplication:
- numerical robust, minimises parameter deviations Δx and circuit strengths δ
- Easy removal of singularities, (nearly) singular eigen-solutions have li~0
  - to remove those solution: if li ≈ 0 → '1/li := 0'
    - discarded eigenvalues corresponds to solution pattern unaffected by the FB
color black{ vec %delta_{ss} = {tilde R} ^{-1} cdot %DELTA {vec x} ~ with ~ {tilde R} ^{-1} = underline V cdot {underline %lambda}^{-1} cdot {underline U}^T ~~ dlrarrow ~~ vec %delta_{ss} = sum from{i=0} to{n} {{a_i} over {%lambda_i} {vec v}_i} ~ with ~ a_i= {vec u}_i^T %DELTA {vec x} }
*G. Golub and C. Reinsch, “Handbook for automatic computation II, Linear Algebra”, Springer, NY, 1971
color white {{underline U}^T underline U = underline 1} color white {underline %lambda = diag(%lambda_1 , .. , %lambda_n) } newline color white {%lambda_1 > %lambda_2 > dotslow > %lambda_n} color white {underline R in setR ^{m times n}} color white {{underline V}^T underline V = underline V {underline V}^T = underline 1 }
U
V
=
T
x
R
x
l
response matrix
BPM eigenvectors
eigenvalues
COD eigenvectors
n x cor. circuits
m x observ-ables

---
Space-Domain:
SVD example: LHC eigenvalue spectrum
Eigenvalue spectra for vertical LHC response matrix using all BPMs and CODs:
near
singular
solutions
dominant eigenvalues
these correspond
to orbit bumps
@ the IPs
condition number ~ 106
→ indicator of matrix condition
→ loss of 12 bits during the inversion process
→ use of 64 bit floats is mandatory

---
Optics and Calibration Uncertainties
Gradient Based Search Features
- Optics imperfections may deteriorate the convergence speed but do not affect absolute convergence (response functions are 'monotonic'):
- Example: 2-dim orbit error surface projection
perfect optic	→ 1 iteration
20% beta-beat	→ ~2 iterations
20% calibration error	→ ~7 iterations