Title
2010-06-23 Internal BI-QP Meeting on SPS BPM Renovation
First Iteration on Requirements and System Qualification
- Motivation – making SPS compatible as high-intensity injector for LHC
  - Timely renovation of ageing MOPOS system
  - Reduction of systematic e.g. fill-pattern dependences
First iteration on SPS position measurement system requirements
- System qualification or 'what are the limits' of the existing infrastructure:
  - Shoe-box, buttons & strip-line pick-ups
  - (Super-long) signal cables
  - Low-level analogue electronic dependencies
  - Digital acquisition System
What are our aims? What can we do?

---
LHC BPM Specification
LHC-BPM-ES-0004 rev. 2.0, EDMS #327557, 2002, p. 25:
Beam threading
Close trajectory on itself
Position error at injection
Momentum mismatch
detection at injection
Optics and local Q' checks
Aperture optimisations
LHC Collimation/Orbit FB
Orbit at injection elements
Position error at injection
Momentum FB (radial loop)
Dispersion measurements
b2/a2 to b4/a4 (~TOTEM)
Are we happy with this type of format? Other constraints? Open...?

---
Early SPS BPM Specification Draft (end-2008, J. Wenninger et al)
Operational beam configurations used/envisaged (Nbatch = 1-13)
- Orbit@1kHz (40 turns) up to 60 seconds, 100 first turns & 10/100k turns total
- Resolution for large intensity beams (> 5·1012 p):
  - orbit mode: 0.1 mm	↔	BPH: 0.1% (& 0.5%) BPV: 0.2%
  - trajectory: 0.4 mm ?	↔	BPH: 0.5% (& 1.8%) BPV: 1.0%
- Resolution for single bunches (LHC pilot):
  - orbit mode: 0.4 mm	↔	BPH: 0.5% (& 1.8%) BPV: 1.0%
  - trajectory: 1.0 mm ?	↔	BPH: 1.3% (& 4.5%) BPV: 2.4%
- Required accuracy: 0.5 mm (w/o alignment) ↔ BPH: 0.6% (& 2.3%) BPV: 1.2
Type	Bunch spacing	Bunch number	Bunch charge [1010]	bunch length [4, ns]
FT / CNGS	5 ns	400-4000	0.1-2	1-4
LHC25NS	24.96 ns	Nbatchx 72	1-20	1-4
LHC50NS	49.92 ns	Nbatchx 36	1-20	1-4
LHC75NS	74.88 ns	Nbatchx 24	1-20	1-4
LHC single bunch	524.4-2022.6 ns	1-16	0.2-20	1-4
LHC ion / Pb82+	100 ns	Nbatchx 4	0.2-2	1-4

---
Future “Operational”/Physically Possible Beam Scenarios
- Operational bunch charges: 1·109 → 2-3·1011 protons/bunch
  - Minimum that can be still accelerated: ~108 protons/bunch
  - Maximum that can be still accelerated: ~2-3·1011 protons/bunch ???
- Time structure: 200 MHz or lower harmonics
    - Single bunch: 1-4 ns
    - Typical batch lengths:
      - SFTPRO/CNGS: ~10 μs (old) vs. ~ 4.4 μs (new)
      - LHC-type beam: down to 1 μs
        - distinction 'batch vs. bunch' less obvious
    - Bunch-by-bunch measurement not strictly required
      - Exception: a few (3-4) pick-ups with LHC-type electronics

---
LSS5 ↔ ECA5 Cable Calibration
LSS4 ↔ ECA4 Cable Calibration
Cable response via open and closed loop S12 setup (NA / pulse generator)
- LSS4 much better bandwidth → not surprising 130 m long 7/8” Heliax cable
- LSS5 more representative w.r.t. cables used by MOPOS: ~ 70 m long CMA50
  - Present MOPOS installation rely on cables that are several 100 m long!
    - Temperature effects, radiation degradation, (replacement) costs, ...

---
Example: ECA5 Measurement BPH vs. BPMB I/III
Raw pick-up signals: BPH (77 mm) vs. BPMB (40 mm half-aperture)
- “Fat” LHC pilot bunch ~2·1010 p/bunch → BPH gives 10x more peak signal
  - N.B. 10dB attenuation for BPMB, 20 dB attenuation for BPH

---
Example: ECA5 Measurement BPH vs. BPMB II/III
Re-aligned and re-normalised signals:
- BPH signals are (by design) shifted by ~0.3 ns	→ any future analogue FE design needs to take this into account, unless re-using the existing (old) cables prior to the hybrid
- BPH has a much more pronounced ringing than the BPMB

---
Example: ECA5 Measurement BPH vs. BPMB III/III
Re-aligned, re-normalised and zoomed-in signals:
- BPH ringing prevails beyond 25 ns → one source of filling-patter dependence
  - worsened by long-cables (bandwidth limitation ↔ dispersion)
  - Transient length and frequency strongly dependent bunch length

---
Ultimate Intensities
Single LHC bunch at 2·1011 p/bunch
- BPH:	~35 V	→ 18 mV /108 protons/bunch
- BPMB:	~ 2.5 V	→ 1.5 mV /108 protons/bunch

---
Typical SFTPRO structure
… problems with second batch during the ramp

---
Terminology: Accuracy & Precision
Good summary: http://en.wikipedia.org/wiki/Accuracy_and_precision
- Accuracy:“[..] closeness of measurements [..] to its actual (true) value”
- Precision (also: reproducibility or repeatability):	“[..] degree to which repeated measurements	under unchanged conditions show the	same results.”
- Example: “Target analogy” and the two extreme cases
Tentative OP specification excludes static alignment errors	→ measurement reference frame is de-facto the BPM body
→ remaining errors reducing the accuracy are drifts/dependences of:	pick-up response, cables, analogue electronics, …
High precision, but low accuracy
High accuracy, but low precision

---
reserve

---
- Some criteria
  - Fast-scale:
    - Turn-by-turn
    - Batch-by-batch
    - Bunch-by-bunch
  - Slow-scale/closed orbit, minimisation of:
    - fill pattern dependencies
    - intensity dependencies
    - Bunch-length dependencies
    - hardware/non-beam dependencies

---
Some Definition
A more formal definition of “Bad”: Distinguish between beam position monitor...
- Error: inconsistency between measured and true beam position
  - minimised by calibration or re-alignment
  - can lead to a a 'Fault' if exceeds pre-defined limits
- Fault or Failure:
  - an error exceeding specified limits or
  - the unavailability of the measurement
N.B.
'accuracy' := maximum measurement error ≠ resolution
'resolution' := minimum measurable position change

---
Sub-Classification of BPM error
In a linear model, errors can be further decomposed into an 'offset and 'calibration factor':
  - (some) errors affect either offset or slope only
  - absolute offset often not required (provided it is constant):
→ e.g. beam-based alignment of LHC Collimation
→ e.g. orbit response or equivalent lattice response measurements
  - systematic calibration factor is minimised by beam-based steering
    - LHC OFB Example: Assuming 20% beta-beat or 20% BPM calibration error, FB reaches after 7 iterations the same convergence as after one iteration for a 1% beta-beat/BPM calibration error.
x_{meas} = x_{offset} ~ + ~ a_{cal} cdot x_{true}