2010-09-17 LHC Machine Committee
Cohabitation of ADT and
Q/Q' Diagnostics Systems
– or –
“Someone's noise is someone-else's signal”
Ralph J. Steinhagen, BE-BI
With input from BE-RF: W. Höfle, D. Valuch et al.

---
Recap:
Q/Q' Diagnostics and Residual Noise on the Beam I/II
- Initial design assumption: no residual tune signatures on the beam (0 dB S/N)
  - Anticipated constant driving of the beam and – to limit the required excitation levels – the highly-sensitive BBQ system was developed
    - further exploited by a FFT and PLL system
  - Hypothesis: BBQ nm-level sensitivity would be sufficient to operate below the “radar” of excitation impacting operation/protection (less than 1 μm)
    - seemed to be confirmed by tests in the SPS, RHIC, Tevatron, …
After the start-up we were blessed (and/or cursed):
  - BBQ proved to provide a turn-by-turn resolution of better than 30 nm
    - 30+dB more sensitivity than other LHC systems (ADT: 1μm, BPM: 50 μm)
  - Ever-present residual Q oscillations on the few 100 nm to few μm level
- Luxurious 30-40 dB signal-to-noise ratios enabled the passive monitoring, tracking and feedbacks without additional excitation
  - reliable from day-one for more than a year now, controlling large tune variations during basically every LHC ramp (and most squeezes)
  - Helped also to identify other beam perturbation issues (mains, hump, etc.)

---
Recap:
Q/Q' Diagnostics and Residual Noise on the Beam II/II
However...
While great for passive monitoring, the nm- to μm-level beam oscillations are incoherent (“noise”) from a FFT/PLL point of view of using explicit excitations.
- Regardless of whether using FFT or PLL:
  - Need to excite ~30 dB above this “noise” to recover the performance of using residual oscillations only (→ 60 dB above BBQ noise floor!):
    - Tune tracking:	min. ~20 dB (assuming |C-|=const)
    - Coupling measurement:	min. 18 dB (better 26 dB)
   → corresponds to ~10 to 100 μm oscillations
  - For comparison: collimators tolerances at about 200 μm
    - tight window between not locking/tracking and causing beam loss
    - uncertainties on BTF due to collective effects, ADT phase/gain, …
  - Driving the beam with such ample signals seemed to be inefficient and less robust compared to the performance achieved with the passive-only system and was considered to be used mainly if the signal would drop...
- Since recently, ADT is used to regularly damp injection oscillations and (with exception of flat-top and squeeze) kept 'on' also during ramp and collisions
  - Damping performance improved from a few hundred turns to < 50!!
  - However: as for any other feedback, higher feedback bandwidths (“gain”) imply also more measurement noise propagated to the beam...

---
ADT Interference on Tune Diagnostic
Example: 0.1 Hz-Avg. BBQ Spectra @450 GeV, one nominal bunch
- BBQ noise-floor raised by 30 dB, wide Q-peak → reduces ΔQres~10-4 → ~10-2
  - Impacts reliable tune (and coupling) measurement & feedback
  - incompatible with Q'-measurements using small Δp/p-modulation
  - loss of additional beam stability diagnostics on mains harmonics, hump, etc.
ADT on (max gain)
ADT off
ADT on (max gain)
ADT off
Beam 1
Beam 2
1 μm
30 nm

---
High-Gain ADT Operation &
Transverse Emittance Growth @ 450 GeV
Not a performance issue: ADT noise/gain does not impact/deteriorate εn
vertical beam size
horizontal beam size
ADT unstable?
Q'>2, ΔQ=-0.006
ADT@max gain
(H: 0, V: -6 dB
ADT@nominal gain
(H/V: -20 dB
1 – 4 mm
1 – 2 mm
~ 12 hours

---
High-Gain ADT Operation &	Transverse Emittance Growth @ 3.5 TeV (50b Physics Fill)
… but has a measurable impact on the achievable tune resolution:
flat-top & sqeeze
ΔQ=0.02
Tunes:
collisions
B2-V
B1 & B2-H
Beam sizes:
B2-V
ADT OFF
~1 hour

---
For comparison:
Residual LHC Tune Stability without ADT
Example (3. ramp 2009-11-30 @00:15):
  - Residual tune stability ΔQ ≈ 5·10-4 at injection energy
    - (in-spec) noise on RQT[D/F] circuits (5mA vs. max. 600 A)

---
Challenge of Measuring Q'(t)
- Real-life test/challenge for required Q-resolution and measurement bandwidth
  - Q'(t)→ ΔQres< ~10-4 @ 2.5 Hz
  - Q(t) → ΔQres< ~10-3 @ 2.5 Hz
Q'(t) via radial modulation (Δp/p=2·10-4@0.25 Hz, limit: res. Q stability @450 GeV)
- With nominal beam (ADT on) “challenging” to measure Q', limits:
  - Δp/p·Q'>ΔQres ~ 0.005 (Δp/p > 2·10-4 impractical/incompatible with nominal beam)
- Default OP procedure: switch ADT 'off' → meas. Q' → switch ADT 'on'
  - Switching ADT 'on'/'off' has little/no impact on lifetime/ε-blowup
Inputs to operators & feedbacks → need to be robust as possible

---
Options to make Q/Q' Diagnostic compatible	with the primary ADT Function I/II
Reduced of tune S/N ratio is primary limiting factor, primary option at hand:
1. Low(er) ADT gain after injection until end-of-squeeze
  1. presently the only viable, reliable and available option until end of 2010
2. High ADT gain for first N-turns after injection, then lower-gain
  1. same as above, but would simplifies operational procedures at injection
Three ADT use-cases affecting the Q/Q' diagnostics:
  - Injection damping (few turns)
  - Damping during collisions	(e.g. beam-beam driven oscillations)
    - very slow tune drifts allow mitigation	via longer averaging periods
  - Operation after injection until end of squeeze → noise is an issue
    - Impact of gain-reduction on day-to-day operation seems to be is small:
      - Little/no impact on emittance growth or beam losses
      - Rare (no?) single- or coupled-bunch instabilities (provided Q'>0)
    - In addition: some operatonal εn-blowup margin in the PS (Mike dicit)
Present situation OK:
no or little impact of high-gain operation on Q/Q' diagnostics

---
Test: ADT Gain/Noise Impact on Q/Q' performance
ΔQ=0.02
6dB gain reduction
- 6dB gain reduction helps but not sufficient for all operational cases (Q', hump, ...)
- Alternative: need to excite the beam... by up to 20 dB more than with ADT 'off'
ADT OFF
ADT OFF

---
Summary:
Options to make Q/Q' Diagnostic compatible with ADT Function II/II
Reduction of tune S/N ratio (30+dB→5dB) is primary limiting factor:
1. Low(er) ADT gain after injection until end-of-squeeze
  1. presently the only viable, reliable and available option until end of 2010
2. High ADT gain for first N-turns after injection, then lower-gain
3. Sacrifical (e.g. non-colliding) bunch for which ADT is disabled/low-gain
  1. ADT ready, BBQ bunch selector needs further development (loss of S/N)
4. Dead-band in ADT gain function masking oscillations below noise floor
  1. Simulation, tests with beam and firmware update required
5. Deriving tune from ADT exciter signal (see additional slides)
  - more operational long(er)-term experience needed w.r.t. robustness, resolution, etc.
1. High ADT gain & Q-PLL exciting ~30+ dB above ADT's noise floor
  1. not without issues: required oscillation amplitudes can reach up to ~100 μm, losses!
  2. complex dependence on ADT gain, energy, intensity, collective effects,...
2. High ADT gain & Q-PLL exciting ~30+ dB above 10x lower ADT noise-floor
  1. same as before, but much preferred as ex. levels are less critical (max. 10 μm)
  2. feasibility of noise reduction needs to be demonstrated
  3. more operational long(er)-term experience needed w.r.t. robustness, etc.
  4. require beam-time for commissioning (e.g. in parallel to regular loss-map checks?)
3. High ADT gain & using tranverse Schottky monitor
  - operational long-term experience needed w.r.t. robustness, achievable bw. etc.

---
Additional supporting slides

---
ADT Dead-Band
- Hypthesis: there are no instabilities that are constantly driving the beam
  - 'True' for present beam configuration but needs revisiting for smaller bunch spacing
- Two different thresholds to control the gain (switch 'off'→ 'on' → 'off')
  - activate damper if instabilities exceed n-um
  - de-activate damper if oscillations are below m-um (e.g. after x-turns)
  - For example: m = 2 μm < n =10-20 um & x = 50
- Strictly: Non-linear hysteresis filter but keeps it linear if ADT is 'on'
- Would fail if frequency of instability occurrences is too high	→ however, should have strong tune signatures in ADT exciter then..
activate
δ
on: 10-20 μm
x
off: 2 μm
de-activate

---
ADT Dead-Band vs. Deriving the Tune from ADT Exciter Signal
- Two complementary options depending on the actual strength and occurrance frequency of instabilities and coupled bunch modes in the LHC:
  - Rare:	→ dead-band is the better option (= damp only unstable beam)
  - Frequent:	→ ADT exciter signal contains modes and their frequencies
    - issue: reliability and achievable meas. bandwidth ΔQres<10-3 @ 2.5 Hz?
vertical tune
courtesy W. Höfle

---
Interdependence between Tune-PLL and ADT
- Tune-PLL not a 'silver bullet' solution but will be further explored:
  - Complex BTF dependence on damper gain/phase, collective effects:
  - Requires excitations 30+ dB above noise floor for reliable signal/lock and coupling measurement: noise ~1 μm → excitation can go up to 100 μm
  - Detected tune peak shifts with effective damper gain: ± 6dB ↔ ΔQ≈3∙10-3
ADT +6dB
ADT default
ADT -6dB
ADT +6dB
ADT default
ADT -6dB

---
Required S/N ratio for Tune and Coupling Diagnostics I/II
- Initial Q-PLL design assumption violated:
  - no residual tune oscillation, need to drive the beam to get some signal
- Non-PLL “random” signals add vectorial to PLL driven one:
  - To lock (ΔQres≈10-4): ~20 dB S/N
  - Once locked: Δφ≈ 0.5° → 8 dB S/N@2.5Hz
    - N.B. un-physical steady-state as Q continuously moving during ramp
matrix{ %sigma(%varphi) # %=% # alignl { arcsin left({%sigma_{f} over A}right)} # %=% # alignl {arcsin left(sqrt{2 over N}{%sigma_{t} over A}right)} ## % # % % approx % # alignl {sqrt{2 over N}{%sigma_{t} over A}} # % # % }
for small noise
to signal ratios
σ(φ)
A
σf
φ
Tune width change →
change of phase slope (K0)

---
Required S/N ratio for Tune and Coupling Diagnostics II/II
- Closest-tune approach not practical while ramping
- Use ratio between regular and cross-term instead:
  - A1,x: eigenmode amplitude '1' in horizontal plane
  - A1,y: eigenmode amplitude '1' in vertical plane
  - requiring resolution so that Δ|C-| < 0.1 |Q1-Q2|, and r = r1=r2 > 0
    → required N/S ration r < ~0.05 ↔ S/N ~ 26 dB
  - requiring resolution so that Δ|C-| < 0.5 |Q1-Q2|, and r = r1=r2 > 0
    → required S/N ~ 20 dB
χ
color red {r_1 = A_{1,y} over A_{1,x}} ~~and~~ color red {r_2 = A_{2,x} over A_{2,y}} newline newline alignl % drarrow ~~~ abs{C^{-%}} = abs{Q_1 - Q_2} cdot {{2 sqrt{r_1 r_2}} over {(1 + r_1 r_2)}} ~~and~~ %DELTA = abs{Q_1 - Q_2} cdot {{(1 - r_1 r_2)} over {(1 + r_1 r_2)}}
R. Jones e.al., “Towards a Robust Phase Locked Loop Tune Feedback System”, DIPAC'05, Lyon, France, 2005

---
Simplified ADT Mechanics
Limit of proportional controller gain and noise
Increasing the gain of D(s) → ∞ (positive only around Q) implies:
 → better attenuation of instabilities but also more propagated noise
Cannot have one without the other…
…requires a trade-off between reducing and δi/δd and minimising the impact of δm
Ext./Int. Excitation
(e.g. AC-dipole, BI input)
Instabilities not governed by the default BTF
~ “scalar gain” K
Beam-Transfer-Function
measurement noise
transfer function/
noise sensitivity:
1
Disturbance rejection:
0
nominal input sensitivity:
0
“A bit less fast”/
“hump damping”

---
The forgotten child: Transverse Schottky
- Operates at a frequency well above (4.8GHz) the ADT bandwidth (<20 MHz)
  - issue: reliability and achievable meas. bandwidth ΔQres<10-3 @ 2.5 Hz

---
Beam Loss in Response to ΔQ=0.005 (Q2.L1)
- Switched from PLL- to k-mod studies (ADT back to nominal, Q'>~2):
  - Missing diagnostics: lost 40% of B2 – ADT saviour or culprit?