2010-10-10 BI Wall-Current-Monitor
BI Wall-Current-Monitor
Description of the
required FESA class functionality
– first iteration –
Ralph J. Steinhagen

---
Definitions and Conventions
- Some definitions:
  - LHC RF frequency 400 MHz
    - 35640 RF buckets
    - 3564 bunch slots, however only 2808 will be nominally filled
- Some conventions:
  - Too many (mostly empty) RF buckets → stick to nom. 25 ns bunch slots and possibly note the nx2.5ns shift if not in nominal bucket
  - Distinguish
    - 'bunch':	requested bunch in nominal RF bucket
    - 'satellite':	undesired bunch elsewhere
      - mostly much below nominal bunch intensity, but
      - could be a nominal bunch with a RF injection bucket error
- Targeted time-scale of observations: few-seconds → hours (via IIR average filter)
- Calculate each parameters per bunch/satellite (arrays) and provide statistic summary of each parameter per beam for quick access/analysis in logging:
  - e.g. '_MEAN', '_MAX', '_MIN', '_STDEV', '_MEDIAN' (?)
  - Example: 'BUNCH_INTENSITY_MEAN', etc....

---
Compensation of WCM System Response
On the Bunch Shape
Real-life bunch does not necessarily obey 'Gaussian' or 'cos²' shapes
- Most difference/details are only visible at very high frequencies > 1 GHz
- Naïve assumption of bunch being shaped according to Gaussian distribution:
  - 1 σ → 68.27% of particles
  - 2 σ → 95.45 % of particles
  - 3 σ → 99.73 % of particles ↔ 0.3% error of intensity estimate (target?)
  - 4 σ → 99.99 % of particles
- Response of pick-up, cables, scope at these frequency need compensation!
Measurement:
Model (multiple-)bunch spectrum:
meas. spectrum
Parab. fit
Gauss-fit
Cos²-fit

---
Compensation of WCM System Response
True longitudinal bunch profile measurement is distorted by:
  1. WCM pick-up response → design values + measurements by T. Bohl &U. Wehrle
  2. combiner-response (star-topology) → only design (re-measure end '10)
  3. Dispersion due to 7/8” Heliax cabling & analogue scope bandwidth
- Accuracies below 10% require compensation (i.e. short bunches)
  - Simple Fourier space deconvolution with measured system response
  - However: (too) high numerical complexity if treating raw 100 us frames
- Propose to:
  - split 100 us frames in 'nx50 ns' and shifted 'nx50 ns + 25 ns' slices
  - make compensation configurable: NONE, FULL, CABLE/SCOPE
after >100 m cable
via optical fiber link
practical/sensible limit
of response compensation

---
Summary of Proposed Beam-Parameter Measurements
- number of bunches & satellites
- bunch length → various estimates:
  - Cos²-Distribution (best?), parabolic-distribution, Gaussian or n x RMS, FWHM, length containing 50/95/99% of power/intensities
  - time-constants of bunch-length increase
  - per-bunch and statistic summary
- long. bunch position → not relevant since we average over seconds
  - shifts w.r.t. nominal bucket position (n x 2.5 ns within 25 ns slot)
  - shift within 2.5 ns bucket
- bunch peak voltage
  - Rise/decay time-constant
  - per-bunch and statistic summary
- bunch intensity of 'bunches' and 'satellites' – nb(B1/2)
  - time-constants of change
  - integrated beam intensity of 'bunches' and 'satellites'
  - per-bunch and statistic summary
- Luminous intensity in the IP – “Σi (nb(B1)*nb(B2))i”
  - IP1, 2, 5 & 8 statistics only	→ rationale: with beam sizes → estimates the machine-lumi life-times
  - time-constants of change

---
Following slides focus more on the technical implementation aspect

---
Compensation of System Response
- Straight deconvolution of the 100 us frame with the system response has probably a (too) high numerical complexity
- Propose to split 100 us frame into 4000 smaller frames,
  - each 50 ns long ↔ 500 samples each (or 512 samples → faster FFT)
  - Small frame start: 10 ns before nominal bunch slot
  - Small frame stop: ~40 ns before nominal bunch slot
- Deconvolution of system response is easiest via Fourier-Transform:
  - Perform FFT of measured frame with 'N'-samples
    - spectrum of containing Re- and Im-component
  - Multiply beam spectrum with selected complex inv. system response:	'NONE', 'FULL', 'CABLE-SCOPE-ONLY'
    - keep intermediate result as it is needed for the intensity and power-based bunch length estimate (too large window, multiple bunches)
    - Will provide the corresponding filters, as a start:	'NONE' ↔ Re(i) = 1 & Im(i) = 0
  - Inverse FFT on compensated spectrum
    - yields compensated frame of 'N' sample length
- For determinism/real-time performance of the FESA server	→ best to perform this for every slot (regardless whether there is a bunch or not)
transient mitigation
Re_{comp} = sum_{i=0}^{N/2} {Re_{data}(i)*Re_{filter}(i)} ~~&~~ Im_{comp} = sum_{i=0}^{N/2} Im_{data}(i)*Im_{filter}(i)

---
Bunch Length Estimates
- Finite Estimates (fit-limits <2.5 ns around peak or 3x noise-floor)
  - COS²-Distribution (probably best):
    - BUNCH_LENGTH_COS2 (DB)/	bunchLengthCOS2 (FESA?)
  - Parabolic-distribution:
    - BUNCH_LENGTH_PARABOLIC/	bunchLengthParabolic
  - 50/95/99% power (by-product of deconvolution/intensity estimate)
    - BUNCH_LENGTH_POWER50,BUNCH_LENGTH_POWER90.../ bunchLengthPower50, bunchLengthPower95,
- Infinite estimates (N.B. non-physical since RF bucket is finite < 2.5 ns)
  - Full-Width-Half-Maximum (see plot):
  - Gaussian distribution:
    - BUNCH_LENGTH_GAUSS/	bunchLengthGauss
  - RMS (alternate to Gaussian)
    - BUNCH_LENGTH_RMS/	bunchLengthRMS
f(t) = I cdot {2 over B} left[cos(%pi{t over B}) right]^2 for ~t in [-B/2, +B/2], 0 ~elsewhere
FWHM = abs{x_2-x_1}
f(x) = 1 over {sqrt{2%pi%sigma_t}} cdot e^{-{1 over 2} left({t-%mu} over {%sigma_t} right)^2}
%sigma = {%<x^2-%<x%>%>~}approx 1 over N sum_{i}^{N} (x-%mu)^2

---
Intensity Estimate and Power-based Bunch-Length I/II
Use system-response compensated spectra, e.g:
- ZOH-interpolation of bunch spectrum to DC
- Intensity nb ~ integral of bunch spectrum up to 2.5-GHz
  - calibration (aka. 'fudge') factor to account for beam-to-pick-up transfer function
  - bunch-length if integral matches, e.g.: σ50% = 1/f if nb(f)/nb(n-GHz) = 0.5,	σ95% = 1/f if nb(f)/nb(n-GHz) = 0.95, ...
to DC

---
Illustration of power-bandwidth
Maximum frequency that contains 50%, 95% and 99% of bunch-spectral power (↔ bunch intensity)
Integrated rel. intensity

---
Rise-/Decay-time Estimates
- → same algorithm as used for the beam-current transformers....
- More news asap.