2013-01-21
LHC Luminosity Calibration & Measurement Working Group
WCM-based Satellite Measurements
during the November'12 VdM scans
Ralph J. Steinhagen, BE-BI
Thomas Bohl, BE-RF

---
Proposed New WCM Cabling Layout (> LS1)
Nov'12 setup
July'12 setup
FXL-780NHR
- For the LHC only one signal path implemented (upgrade as for PS?)
- All transfer function (except WCM) calibrated and under control down to <10-4

---
2012: First Beam Data after Modifications –
July vs. November VdM Scan B1
- During last TS: synchronised setup B1 & B2:	simplified setup → reduced systematics by removing star-combiner
- re-measured cable/scope transfer function → eliminates all reflections except those created within the WCM itself (N.B. cannot back-propagate these changes for B2 measurements during VdM scan in November)
original setup
w/o star-combiner
confirmed
satellite

---
Typical WCM Signal Response – Linear Scale
Removed the star-combiner for B1 (source for sys. reflection after bunch)
WCM B1
full scale
1st bunch
WCM B2
full scale
1st bunch
±1% zoom
1st bunch
±1% zoom
1st bunch

---
Typical WCM Signal Response – Log Scale
2.5 ns
10 ns ??
1% level
-5 ns
-30 ns

---
Typical WCM Signal Response – Log Scale - Zoom
2.5 & 5ns ns visible
but on large BG
(hard to quantify)
obvious satellite @- 5 ns

---
VdM Scan 21th November – B1 I/II
Superimposed Raw Traces

---
VdM Scan 21th November – B1 II/II
Integrated over RF Bucket

---
VdM Scan 21th November – B2 I/II
Superimposed Raw Traces

---
VdM Scan 21th November – B2 II/II
Integrated over RF Bucket

---
VdM Scan 23th November – B1 I/II
Superimposed Raw Traces

---
VdM Scan 23th November – B1 I/II
Superimposed Raw Traces

---
VdM Scan 23th November – B2 I/II
Superimposed Raw Traces

---
VdM Scan 23th November – B2 I/II
Superimposed Raw Traces

---
VdM Scan 21th November – Satellite Summary
- Beam 1:
  - @-30 ns: nb ≈ 0.1 ... 0.44·10-3 <Ib>max	→ total: 0.17·10-3 IB1
  - @- 5 ns: nb ≈ 0.3 … 0.6·10-3 <Ib>max	→ total: 0.31·10-3 IB1
  - @+2.5 ns: nb ≈ 3.0 … 4.0·10-3 <Ib>max	→ total: ??? IB1
  - @+10 ns: nb ≈ 0.2 ·10-3 <Ib>max	→ total: 0.6·10-3 IB1 <not confirmed>
- Beam 2:
  - @-30 ns: nb ≈ 0.1 ... 0.3·10-3 <Ib>max	→ total: 0.15·10-3 IB1
  - @- 5 ns: nb ≈ 0.2 … 0.8·10-3 <Ib>max	→ total: 0.35·10-3 IB1
  - @+2.5 ns: nb ≈ ??? <likely there but dominated by systematic>
  - @+35/45 ns: nb ≈ small <Ib>max	→ total: 0.6·10-3 IB1 <not confirmed>
B1 & B2: no sig. contributions outside [-100 ns,+100ns] w.r.t. main bunch

---
VdM Scan 23th November – Satellite Summary
- Beam 1:
  - ? bunches@-50 ns: nb ≈ ???·10-3 <Ib>max	→ total: 0.21·10-3 IB1
  - 4 bunches@-30 ns: nb ≈ 0.2 ... 0.44·10-3 <Ib>max	→ total: 0.232·10-3 IB1
  - 7 bunches@-5 ns: nb ≈ 0.2 … 1.3·10-3 <Ib>max	→ total: 0.382·10-3 IB1
  - ? bunches@+2.5 ns: nb ≈ 3.0 … 4.0·10-3 <Ib>max → total: ??? IB1
  - 7 bunches@+10 ns: nb ≈ 0.2 … 1.3·10-3 <Ib>max → total: 1·10-3 IB1
- Beam 2:
  - 3 bunches@-30 ns: nb ≈ 0.3 ... 0.4·10-3 <Ib>max → total: 0.18·10-3 IB1
  - 5 bunches@- 5 ns: nb ≈ 0.2 … 1.0·10-3 <Ib>max → total: 0.45·10-3 IB1
  - @+2.5 ns: nb ≈ ??? <likely there but dominated by systematic>
  - @+35/40/45 ns: nb ≈ small <Ib>max → total: ~10-2 IB1 <not confirmed>
B1 & B2: no sig. contributions outside [-100 ns,+100ns] w.r.t. main bunch

---
Summary
Improved WCM systematics for B1 and B2 (after Nov'12)
- WCM resolution ±~5·10-5(r.m.s.) ± ~6·10-5(sys.)for an integration period of ~20 s
- Satellite detections systematic:
  - t= 0-15 ns: Isat(sys)<5·10-3 <Ibunch>max, otherwise: Isat(sys)<10-5 <Ibunch>max
N.B. presented satellite numbers based on following cut threshold
  1. before main bunch:
    - raw signal > 2·10-4 w.r.t. <Ibunch>max or
    - Integrated signal > 2·10-4 w.r.t. <Ibeam>max
  1. After main bunch:
    - > 20 ns as for 'I'
    - < 20 ns not reproducible w.r.t. bunch-to-bunch	(i.e. not a reflection, introduces bias for satellites that are similar from b-2-b)
Residual error given by reflections caused within WCM	→ planned to be re-measure during LS-O